# Non-Conformance Register

Log of Non-Conforming Work & Corrective Actions

Year: 2024
Lab: [Lab Name]

This register tracks all identified non-conformities (NCs) from audits, complaints, or daily operations.

| ID # | Date | Source | Description of Issue | Risk Level | Status | Closed Date |
| --- | --- | --- | --- | --- | --- | --- |
| NC-24-01 | Jan 10 | Internal Audit | Balance #4 calibration expired. | High | CLOSED | Jan 15 |
| NC-24-02 | Feb 02 | Cust. Complaint | Report issued with wrong unit. | Medium | OPEN |  |

# Corrective Action Report (CAR)

Root Cause Analysis & Improvement

CAR ID: NC-24-02
Date: Feb 02, 2024

## 1. Description of Non-Conformity

Identified By:

[Name / Role]

Detailed Description:

The final report for Client X showed results in mg/L, but the method requires ug/L. This caused a confusion in compliance reporting.

Immediate Correction:

1. Report withdrawn immediately.
2. Amended report issued with correct units.
3. Client notified via email.

## 2. Root Cause Analysis (RCA)

Use 5 Whys or Fishbone method to find the true cause.

| Why? (1) | The analyst selected the wrong template in LIMS. |
| --- | --- |
| Why? (2) | Both "Water" and "Waste Water" templates look identical in the dropdown menu. |
| Why? (3) | Naming convention is "Template A" vs "Template B" - not descriptive. |
| Root Cause | Poor naming convention in LIMS software makes it easy to select the wrong unit template. |

## 3. Corrective Action Plan

| Action Steps | Responsible | Due Date |
| --- | --- | --- |
| Rename LIMS templates to clearly state units (e.g., "Water - mg/L"). | IT Manager | Feb 10 |
| Train staff on new naming convention. | Quality Mgr | Feb 12 |

## 4. Verification of Effectiveness (Closure)

Action Taken?

Yes, templates renamed on Feb 10.

Evidence of Success:

Checked 10 reports generated after the change. 100% used the correct template. Error has not recurred.

Quality Manager Approval: [Sign Here]

Date Closed: [Date]